NB-532-3-B
Narrow bandpass filter, 532 nm, 3.0 nm FWHM, 25.0 mm mounted

Sespe Optics
Thin-film filters and precision optics

Datasheet NB-DS-532-3-B, Rev B, 2026-10-24. Replaces Rev A of 2026-06-12.

The NB-532-3-B passes a 3.0 nm band centered on 532 nm and blocks 400 to 1100 nm outside CWL ± 12 nm to at least OD 5. It is a three-cavity, ion-beam-sputtered, all-dielectric coating on a UV fused silica substrate, with a broadband antireflection coating on the rear face, mounted in a black anodized aluminum ring. The blocking is reflective: there is no laminated absorbing glass in the stack. Center wavelength, bandwidth and peak transmission are measured on every filter.

| Center wavelength | Bandwidth (FWHM) | Peak transmission | Blocking |
| --- | --- | --- | --- |
| 532 nm | 3.0 nm typ | 85 % min | OD 5 min |
| 531 to 533 nm at AOI 0°, 23 °C | 2.6 to 3.4 nm at AOI 0°, 23 °C | 90 % typical, at the measured CWL, AOI 0° | 400 to 1100 nm, outside CWL ± 12 nm; OD 6 typical |

# Features

- Center wavelength 532.0 nm, ±1.0 nm at AOI 0°, 23 °C

- Flat-top passband, 3.0 nm FWHM typical

- Peak transmission 90 % typical at the center wavelength, AOI 0°

- Blocking OD 5 minimum, 400 to 1100 nm outside CWL ± 12 nm

- Operates from −20 to +70 °C; stable after humidity exposure (lot sample)

# Applications

- Laser-line cleanup for 532 nm Raman and fluorescence excitation

- Narrow receiver filter for 532 nm lidar and ranging

- Laser-line imaging and beam profiling with a green source

- Background rejection in front of photodetectors and spectrometers

# Ordering information

Part number key: NB - 532 - 3 - B. The limits in this datasheet are those of NB-532-3-B; the other sizes of the same part carry the limits in Table 1.

Table 1. Part number key and sizes of the 532 nm, 3 nm part

| Part number | Size | Mount diameter | Clear aperture | CWL limits | FWHM limits |
| --- | --- | --- | --- | --- | --- |
| NB-532-3-A | size A | 12.5 mm | 9.0 mm | 531.0 to 533.0 nm | 2.6 to 3.4 nm |
| NB-532-3-B | size B (this datasheet) | 25.0 mm | 21.0 mm | 531.0 to 533.0 nm | 2.6 to 3.4 nm |
| NB-532-3-C | size C | 50.0 mm | 45.0 mm | 530.5 to 533.5 nm | 2.5 to 3.5 nm |

Key fields: NB is the series (narrow bandpass, hard-coated, mounted); 532 is the center wavelength in nm (series options 405, 450, 488, 532, 561, 633, 780, 808, 852, 980, 1064); 3 is the nominal FWHM in nm (1, 3 or 10); B is the size, A 12.5 mm, B 25.0 mm or C 50.0 mm mounted.

# Test policy

Every unit: center wavelength, bandwidth and peak transmission at AOI 0°, 23 °C, with the results supplied with the filter. A sample from each coating lot: blocking from 400 to 1100 nm, and humidity and temperature cycling to MIL-STD-810 style methods (tested to, not certified), with no measurable center wavelength shift after 24 h at 60 °C and 85 % relative humidity. Typical values and curves are not tested per unit.

# Absolute maximum ratings

Table 2. Limits that must not be exceeded. They are kept apart from the characteristics in Table 3.

| Parameter | Conditions | Min | Typ | Max | Unit |
| --- | --- | --- | --- | --- | --- |
| Operating temperature | filter in its mount | −20 |  | +70 | °C |
| Storage temperature | dry, unmounted or mounted | −40 |  | +85 | °C |

# Optical characteristics

Table 3. Optical characteristics of NB-532-3-B. Minimum and maximum are tested limits; typical values are for information and are not tested on every unit (see Test policy).

| Parameter | Conditions | Min | Typ | Max | Unit |
| --- | --- | --- | --- | --- | --- |
| Passband |  |  |  |  |  |
| Center wavelength, CWL | AOI 0°, 23 °C, unpolarized, collimated 2 mm beam at the clear aperture center | 531 | 532 | 533 | nm |
| Bandwidth, FWHM | AOI 0°, 23 °C | 2.6 | 3.0 | 3.4 | nm |
| Peak transmission | at the measured center wavelength λ0, AOI 0° | 85 | 90 |  | % |
| Attenuation at CWL ± 5 nm | AOI 0°, typical curve |  | 3 |  | OD |
| Attenuation at CWL ± 12 nm | AOI 0°, typical curve |  | 5 |  | OD |
| Blocking |  |  |  |  |  |
| Blocking | 400 to 1100 nm, outside CWL ± 12 nm, AOI 0° | 5 | 6 |  | OD |
| Blocking test limit | test dynamic range; higher readings reported as > limit |  |  | 6.5 | OD |
| Rear surface reflectance | broadband AR coating, average 400 to 1100 nm |  |  | 0.5 | % |
| Angle and temperature |  |  |  |  |  |
| Effective index, neff | average of s and p, fitted at 10° and 20° | 1.9 | 2.0 | 2.1 |  |
| CWL shift at 10° AOI | typical, from neff = 2.0, Eq. (1) |  | −2.0 |  | nm |
| CWL temperature coefficient | 23 to 70 °C, AOI 0° |  | +4 | +6 | pm/K |
| Laser damage |  |  |  |  |  |
| Damage threshold, pulsed | 532 nm, 10 ns, 20 Hz |  | 0.5 |  | J/cm2 |
| Damage threshold, CW | 532 nm, continuous wave |  | 1 |  | kW/cm2 |

Attenuation at CWL ± 5 nm and ± 12 nm is typical, not a limit; the typical curve in Figure 1 gives OD 3.2 and OD 5.4. Blocking is stated outside CWL ± 12 nm because the passband skirts are still falling inside it. Damage thresholds are typical and not tested on shipped units.

# Tolerance notes

The center wavelength limit of 531 to 533 nm is for AOI 0° at 23 °C and temperature adds to it. At the maximum coefficient (+6 pm/K), and taking it to apply below 23 °C as well, the guaranteed window over −20 to +70 °C is 530.74 to 533.28 nm; the typical drift over that range is 0.36 nm.

Polarization. Limits apply to unpolarized light at AOI 0°. In a representative three-cavity coating of this type the s and p center wavelengths differ by 0.07 nm at 20° and 1.6 nm at 45°; ask for s and p curves for polarized beams above 20°.

# Typical curves

Figure 1. Typical transmission of NB-532-3-B, AOI 0°, 23 °C, unpolarized. (a) Passband on a linear scale, 500 to 560 nm. (b) Blocking as optical density, 400 to 1100 nm. Typical curve from the passband model of the NB series (three-cavity shape, FWHM 3.0 nm, peak 90 %, blocking floor OD 6); it is not a measurement of one filter.

Figure 2. Center wavelength against angle of incidence in air, unpolarized, from Eq. (1) with neff of 1.9, 2.0 and 2.1. The dotted line marks 10°.

Figure 3. Center wavelength against temperature, AOI 0°, typical coefficient (+4 pm/K) and maximum (+6 pm/K), both through 532.0 nm at 23 °C.

# Using the filter at an angle

Tilting moves the passband to shorter wavelengths. In air, with neff from Table 3,

|  | λ(θ)=λ01−sinθneff2 | (1) |
| --- | --- | --- |

For neff = 2.0 the center wavelength is 529.99 nm at 10°, a shift of −2.01 nm, or 67 % of the bandwidth. The shift is −0.51 nm at 5°, −7.84 nm at 20° and −34.36 nm at 45°. At 10° it is −2.23 nm for neff = 1.9 and −1.82 nm for 2.1.

Tilt tuning. Tilt only lowers the center wavelength. A filter that measures 533.0 nm at AOI 0° reaches 532.0 nm at 7.0° (6.7° to 7.4° for neff from 1.9 to 2.1). A filter at the 531.0 nm lower limit cannot be tuned up to 532 nm.

Converging and diverging beams. Each ray sees its own angle. In an f/2 beam (half angle 14.04°) the pupil-averaged shift is −2.00 nm at normal incidence and −3.93 nm at 10° tilt, and the passband broadens. Use the filter in collimated space where bandwidth matters. The calculator at jezdangomez.com/calculators/aoi-shift covers other beams.

# Mechanical

Figure 4. Outline of NB-532-3-B (size B), dimensions in mm. Front view with the element shown dashed, and section A-A through the upper wall at a larger scale. The formal drawing is NB-25-B-OD-001, Rev A.

Table 4. Mechanical and material characteristics of NB-532-3-B

| Parameter | Value | Tolerance or limit | Unit |
| --- | --- | --- | --- |
| Mounted outside diameter | 25.0 | +0 / −0.1 | mm |
| Mounted thickness | 5.0 | ±0.1 | mm |
| Clear aperture, diameter | 21.0 | min | mm |
| Substrate, UV fused silica, dia × thick | 22.0 × 2.0 | reference | mm |
| Transmitted wavefront, 633 nm, over the clear aperture | λ/4 | max |  |
| Wedge | 3 | max | arcmin |
| Surface quality | 60-40 | scratch-dig |  |
| Coating and mount | ion-beam-sputtered Nb2O5 / SiO2 multilayer on one substrate; black anodized aluminum ring, arrow engraved |  |  |

# Handling and mounting

- The arrow engraved on the ring points in the direction of transmission. The coated face, on the lip side of the ring, faces the light source.

- Clamp on the ring only; do not load the element. Maximum retaining ring torque 0.15 N·m.

- Handle by the ring. Blow off dust with dry nitrogen; clean with spectroscopic-grade alcohol on lens tissue; no ultrasonic bath. Store dry between −40 and +85 °C.

# Revision history

Rev A, 2026-06-12. Initial release: four values and one curve.

Rev B, 2026-10-24. Test conditions added to every specification; typical values separated from limits; blocking stated outside CWL ± 12 nm; three curves drawn from code; tolerance notes, test policy, part number key, outline and handling notes added.

Sespe Optics is a fictional company. Portfolio sample by Jezdan Gomez, jezdangomez.com.